PROJECT SUMMARY
Project Title
Building High-Quality, Low-Cost Medical Devices Using 3D Printers and Rapid Prototyping Technology

PLAIN LANGUAGE ABSTRACT
Please describe the proposed project using non-technical language emphasizing the research question and importance to health care delivery. Max 1250 characters
The Glia Academic Collaboration is a medical research project whose goal is to develop, validate, certify and disseminate high-quality, low-cost, open-access medical devices that have been approved by Health Canada. These open source devices will be available for wide use in Canada and internationally, especially at low-volume hospitals and low-income communities.
The availability of inexpensive and modifiable medical devices will improve patient care by giving rural and disadvantaged communities better access to important, life-saving, quality medical devices. Devices like this are currently unattainable due to high costs or logistical barriers.
The project philosophy allows for open access to all work, data and procedures, as well as access to software and hardware plans of finished products. These efforts will enable local, national and international hospitals and ministries to build their own high-quality, safe and affordable medical devices. This allows for better supply chain management, reduces repair costs, and minimizes time during which equipment is inoperable. This will result in significant cost savings for clinics, hospitals and ministries of health.
What is the innovation and main question or hypothesis for this proposal? Max 1250 characters
The Glia Academic Collaboration's innovation relies on two pillars: the use of 3D printers and other rapid prototyping technology; and leveraging Open Access and Open Source principles and devices to decrease development costs and disseminate results to stakeholders.
The introduction of these pillars to traditional device development models allows rapid and inexpensive development of critical devices and wide distribution to low-resource or low-availability settings. In addition to lowering the cost of entry, devices can be modified for local conditions. These variants can easily be contributed to publicly available repositories for other centres to evaluate and use.
This model has been proved with a simple medical device (the stethoscope), and is currently being extended to devices such as the pulse oximeter, electrocardiogram, and hemodialysis.
Please describe how this project will influence the health care system or influence patient care. Max 1250 characters
Our project will be influential to small clinics and underserviced hospitals as they will be able to own and maintain high-quality medical devices with decreased cost burden. Our project will aid hospitals and ministries of health in the direct manufacture of this device. The availability of research-validated open source designs will also allow current low-cost manufacturers to re-enter the market with higher quality devices and will exert overall downward pressure on commodity medical device costs. The availability of high quality generic brands will in turn exert pressure on premium brand manufacturers to match or reduce prices. This has been observed in the consumer and pharmaceutical spaces, and there is no barrier to its replication in the medical device space.
In addition, any team that wishes to modify the device for their needs can do so without having to expend time and money to reverse engineer consumer devices, license devices or gain buy-in from manufacturers.
One societal impact of this work is to allow equal-access to a gold-standard medical device for patients in poor or disadvantaged communities.

PROJECT NARRATIVE PROPOSAL
Background, Objectives, Rationale, Significance, and Justify why the project is innovative. Also include an assessment of expected outcomes. Max 8000 characters
Background and Rationale
The high cost of modern medical diagnostic and therapeutic devices renders them inaccessible to solo clinics, junior physicians, students, nurses and other allied professionals. In some low-resource institutions such as in rural clinics, LHINs and counties, there may be fewer devices than are needed. Sometimes, devices are even completely unavailable. This results in a reduction of the standard of health care compared to richer and more urban institutions. In Ontario, this division impacts populations that are more rural, more poor, more aboriginal and more coloured disproportionately, perpetuating cycles of poverty and disenfranchisement.
The Glia Project uses an open access research, development and distribution model to create high-quality low-cost medical devices that are then clinically validated. This allows off-patent devices to exist as high-quality generic models, which will eventually exert downward pressure on prices for high-quality premium brands.
The stethoscope is a mature proof-of-concept for this model. Stethoscopes are simple instruments that have undergone little innovation in design or fabrication for decades. Despite this, the Littmann Cardiology III, widely regarded as the field’s gold standard, still costs several hundred dollars. This cost is a barrier in Ontario to allied professionals, and especially strikes volunteer medical providers in rural parts of Ontario.

Modeling off the 1963 Littmann design whose patent expired decades ago, we created a stethoscope that performs on par with the Cardiology III (using a chest wall phantom and analyzing spectral analysis between models) at a cost of $2.83 USD (images and supporting data included in Appendix A). To reduce cost, 3D printers are used to create several parts including the stethoscope head, Y-piece, spring and ear tubes. This device has been validated and is certified by Health Canada as a class I device.
Pulse oximetry is also a vital tool in the rapid and non-invasive assessment of patients, providing a continuous estimate of hemoglobin oxygen saturation in arterial blood. Unfortunately, the costs of these devices are prohibitive and reduce availability in smaller centres, putting patients in danger of easily treatable and preventable conditions. With current rapid prototyping technologies such as 3D printing, it is possible to create a very inexpensive pulse oximeter that meets or exceeds the gold standard.
Objectives
The goal of this project is to develop, validate and certify a pulse oximeter that measures hemoglobin, carboxyhemoglobin and methemoglobin. This pulse oximeter will be certified with Health Canada, and then released under the Open Hardware License (OHL), such that hospitals and LHINs in rural and impoverished communities in Ontario would have easy access to these devices.
The project reduces the asymmetry of care by providing physicians and allied professionals – especially those in low-income parts of the province – with access to high-quality, low-cost medical equipment.
To do this, we will:
- Complete engineering prototypes and make modifications during the project as needed.
- Complete a calibration and validation study of our 3D printed pulse oximeter prototype with approximately 400 participants (50 calibration; 350 validation).
- Gain regulatory approval through Health Canada.
- Distribute open source pulse oximeters to hospitals, ambulances and clinic in the LHIN.
- Create partnerships with hospital biomedical departments to manufacture and distribute pulse oximeters in London and the remainder of our LHIN to nurses, paramedics, medical students, junior doctors and other allied professionals.
- Publish all specifications, production files and data in a way that is unencumbered by patents or other restrictions on reproduction (Open Source).

Significance
This project will impact the availability and quality of pulse oximeters in Ontario, especially in solo clinics, rural hospitals and remote outposts. It will also encourage the repair of devices onsite, reducing equipment downtime and repair costs.
Glia's stethoscope has already made a significant impact, being inducted into the Canada Science and Technology Museum and being covered by numerous news publications. It has also received interest from Canadian medical students and allied professionals and is being distributed in London, Ontario followed by other Ontario centres with limited access to medical devices. There are even international hospitals in low-resource settings that are self-sufficient producers of these stethoscopes using recycled plastic. Several staff in the Division of Emergency Medicine already use these stethoscopes.
The Glia Project team produced approximately 200 stethoscopes and distributed them to Western University staff, students and residents. Our team is producing another 1000 stethoscopes for the remainder of the Division of Emergency Medicine. The final design is openly available for production by other medical schools and hospitals.
Replicating this success with pulse oximeters will lead to the wide availability of devices. The availability of specifications for generic manufacturers to manufacture devices and the subsequent downward price pressure on premium brand manufacturers will increase the standard of care for all patients in Ontario. It will also allow LHINs to save costs while maintaining equivalent quality of care.

Innovation
The Glia Project innovates in several major domains: The use of rapid prototyping technology such as 3D printing; The use of Open Source Hardware and Software methodologies; The Open Access approach to distribution of devices and design files.
Rapid prototyping technologies such as 3D printing are common in the research and development stages of medical device creation. However, no market devices exist that are mass-produced using these technologies. Our team has made significant strides in eliminating problems associated with the use of 3D printing for mass manufacture, and will continue to develop this versatile and adaptable technology for mass manufacturing. Our team is at the leading edge of this technology.
Open Source Software development started under the GNU Public License in the 1980's. In the past decade, this methodology has been applied to a nascent DIY electronics movement. As yet, there are few makers of Open Source Hardware, and none that make pulse oximeters that can be safely and legally used by health care providers. This project combines the techniques of Open Source Hardware and Software development with the highly regulated field of medical devices to create a brand new space in biomedical engineering.
Open Access publication of journal articles has been popular for at least the past decade since the inception of the Public Library of Science (PLoS). However, the label seldom encompasses what the Glia Project has released: Complete bill of materials, schematics, validation data and aids to recreate and test devices. To the best of our knowledge, no certified medical device has ever done this.
The Glia Project's innovations will allow medical institutions and medical schools in Ontario to manufacture and repair their own devices for medical students and allied professionals as needed. Each hospital or medical school could create an inexpensive but state-of-the-art facility where production, distribution and further innovation can take place. Any medical device the project develops subsequently can join the repository of devices available for production by these centres.

Assessment of expected outcomes
Our outcome measures are as follows:
- Number of total deployed pulse oximeters
- Number of participating medical schools
- Number of participating hospitals
- Number of participating health care centres (non-hospital)
- Calibration study of pulse oximeter (yes/no)
- Validation of pulse oximeter (yes/no)
- Health Canada certification of pulse oximeter (yes/no)
Team – Please include the names and affilations of the team members, with a brief statement of their roles, experience, area(s) of expertise, time commitment and planned contributions to the project.
Max 2000 characters.
Tarek Loubani, Primary Investigator
Dr. Loubani is an Associate Professor in the Division of Emergency Medicine, Department of Medicine. He has been involved in technological and engineering research for over a decade, spending the past 4 years researching cost reductions of common medical devices. His role is as PI, clinical expert and engineering liaison. He will work on the project three days per week.
Carrie Wakem, Project Manager
Mrs. Wakem is the full-time Project Manager. She supported the project on an administrative level for the past 3 years. Her responsibilities will include: managing the team financial and evaluation reports, maintaining a publicly accessible log of the progress of the project, organizing weekly team meetings, project taskboard, calendar, quarterly reports, and legal and financial paperwork.
Melanie Columbus, Research Coordinator
Dr. Columbus is a full-time research coordinator with the Division. She has assisted in the design of the study and will further assist during data collection, analysis and publication.
Alex Pavlosky, Research Assistant
Mr. Pavlosky is a third year medical student at Western University with the project since 2016 through the SRTP Program. He assisted with the design and publication of the stethoscope research. He will be helping with the pulse oximeter clinical trials and publication preparation.
Luka Mustafa, Engineer, Inštitut za razvoj naprednih aplikativnih sistemov Rače (IRNAS)
Mr. Mustafa is the Chief Operating Officer of IRNAS, based in Slovenia and PhD candidate in engineering at University College London. IRNAS is an internationally renowned open-hardware and electronics engineering firm, with experience from sound-system to particle accelerator control systems. IRNAS is engineering the Pulse Oximeter prototypes and will iterate the design as necessary to enable high quality and mass production. They will also obtain ISO 1394 certification to meet regulatory obligations.
Approach and Work Plan – Provide a description of the approach to be used as well as a detailed project work plan. Max 4000 characters.

Phase I – Device Design & Prototype Creation
We conducted significant research using publicly available documents, expired patents, previous published research and hobbyist efforts found online. Due to device complexity, this process was led by biomedical engineers from the IRNAS Institute with heavy input from the Glia Project team for clinical specifications.
Rapid prototyping techniques such as 3D printing, CNC routing and laser-cutting allowed many iterations of our devices in a short period of time. Our first pulse oximeter prototype is now ready for phase II.
Phase II – Experimental Calibration & Validation (REB approved)
All study procedures will be conducted in the emergency at LHSC with an emergency-trained physician (EP) present at all times.
Calibration will involve approximately 50 participants. Participants will be cannulated with an indwelling arterial catheter for frequent sampling of arterial blood. The experimental pulse oximeter and a gold-standard pulse oximeter will be placed on the participant's finger and continuously monitored. Participant samples will be tested with a CO-oximeter to ensure that background carboxyhemoglobin (COHb) is less than 2%. Under careful monitoring by the EP, oxygen saturation level will be decreased with gas mixtures of progressively decreasing oxygen and increasing nitrogen. Measurements will be recorded until the participant has unstable vitals (e.g., hypotension, tachycardia), is symptomatic or a pulse saturation of 70% is achieved. Oxygen levels will be immediately restored. Arterial blood samples will be drawn and oxygen saturation will be measured when the gas mixture and R value readings have stabilized. 2 cubic centimetres of blood will be drawn for every 5% drop in blood oxygen saturation from 100% to 70% or earlier if the participant experiences unstable vitals or is symptomatic. This part of the study will take approximately 60 minutes to complete. This study design is routine for all pulse oximeters.
Validation will involve approximately 350 community volunteers and ED patients. Investigators will approach patients at the invitation of a member in the circle of care to discuss the study. After going through the LOI and gaining documented written consent, the volunteer/patient will be enrolled. The experimental and control pulse oximeters will be placed on two different fingers on the same hand. Patients will then have one measurement taken with the experimental pulse oximeter and another measurement taken with a gold standard pulse oximeter. Measurements of oxygen saturation will be recorded to establish equivalence of the two oximeters.
Phase III – Health Canada Approval and Publication
The pulse oximeter must pass regulatory approval to be ethically used on patients in Ontario hospitals such as LHSC. This additional layer of scrutiny will ensure the safety and effectiveness of the device.
All designs, protocols and data will be shared openly with the medical community via open access journals. All source data will be deposited in an open repository made available by the Glia Project team.
Phase IV – Assembly Instruction & Dissemination of Technology and Information
Instructions to assemble and test the device will be prepared via video and written instruction.
We will help interested medical institutions (such as LHSC), medical schools and LHINs to safely manufacture and deploy devices. This will ensure broad participation and constant collaboration with other stakeholders and users.
Timeline – Provide a target start date, the dates when each milestone will be completed and the target finishing date for the project. Max 1300 characters
The Glia Project team has completed Phase I of our pulse oximeter project and will enter Phase II in early 2018. Phase II will consist of a calibration trial of approximately 50 healthy volunteers. A validation trial will follow in the second quarter of 2018. The London Health Sciences Centre Emergency Department will be the host for these trials.
After device calibration and validation, the team will seek Health Canada approval starting in the last quarter of 2018. The Health Canada approval is expected to take approximately nine months to complete. During this time we will also prepare and seek publication in a leading medical journal.
When regulatory approval is obtained, we will collaborate with LHSC and other hospitals and LHINs to manufacture and deploy the devices for practical evaluation and feedback. This will be followed by wider roll-out. This phase will begin in the second quarter of 2019.
Project Outcome and Evaluation Metrics – Detail the anticipated outcomes and how these outcomes will be measured. Which measurement will determine success? How will you measure it? What is the expected benefit and how will it be measured? How will the IFPOC determine if the project was successful? Max 2700 characters
The primary outcome of this project is to produce a safe, high-quality, low-cost, open source pulse oximeter that is accessible to all health care providers in Ontario.
This outcome can be measured through the following:
- Number of total deployed pulse oximeters
- Number of participating medical schools
- Number of participating hospitals
- Number of participating health care centres (non-hospital)
- Calibration study of pulse oximeter completed (yes/no)
- Validation of pulse oximeter completed (yes/no)
- Health Canada certification of pulse oximeter obtained (yes/no)
- Publication in at least one open-access medical journal (yes/no)

The expected benefits will break down into three categories:
- Reduced cost to hospitals and non-hospital centres
- Increased availability of pulse oximeter devices
- Decreased device downtime and turnaround

In this project, we intend to measure financial benefit by examining the change in procurement budget over the time our pulse oximeters are being used to replace premium brand devices. Availability would be a straightforward comparison of before and after device numbers. We cannot currently effectively measure device downtime.

Success will be achieved if the following milestones are reached:
- Completion of calibration and validation studies
- Regulatory approval from Health Canada
- Deployment of at least 15 pulse oximeters within LHSC
- Deployment of at least 3 pulse oximeters at three other hospitals in southwestern Ontario
- Deployment of at least 2 pulse oximeters at three non-hospital medical centres in southwestern Ontario

Additional Project Funding – Elaborate on any supplemental funding for the proposed project including the amount and source as well as any contingencies or other special circumstances. Max 2000 characters
The project receives funding support from a prestigious fellowship from the Shuttleworth Foundation. Dr. Tarek Loubani became a fellow of the foundation in March 2017, and is eligible for three one-year terms renewable each March.
Funds from the Shuttleworth Foundation are allocated as follows:
1. Salary support for Dr. Tarek Loubani of up to $75,000 USD per annum.
2. Project matched funding of up to $250,000 USD if Dr. Tarek Loubani forfeits $25,000 USD of his salary support for 10:1 matched funding from the foundation. These funds do not have a fixed expiry date.
In addition, Dr. Loubani has donated a significant amount of personal funds to the project totaling over $200,000 CDN so far in this project, including approximately $50,000 to the pulse oximeter work so far. This is a clear signal of commitment to the project and its longevity. These funds are donated, and there is no mechanism by which this funding will be returned to Dr. Loubani.
Other Relevant Information – Please elaborate on any other items you think are important or relevant to the implementation or outcomes of this project. Max 2000 characters

List references and/or support from partners Max 3000 characters
The Division of Emergency Medicine is committed to this project and believes in its ambitious goals to reduce the disparity to access of high-quality medical devices. The Division has agreed in principle to use devices in the ER that we develop. The Division will also allow us to use space at Victoria Hospital in London during the clinical trial of the pulse oximeter. This will include both the calibration and validation of the device. The same will be true of future devices.

# References
1.	Walker, I. & Wilson, I. Anaesthesia in developing countries - a risk for patients. Lancet 968–969 (2009).
2.	Haynes, A., Weiser, T. & Berry, W. A surgical safety checklist to reduce morbidity and mortality in a global population. New Eng J Med 491–499 (2009).
3.	Duke, T., Graham, S. & Cherian, M. Oxygen is an essential medicine: a call for international action. Int J Tuberc Lung Dis 1362–1368 (2010).
4.	Ashraf, H., Jobayer, M. & Alam, N. Treatment of childhood pneumonia in developing countries. (2010).
5.	Kwizera, A. & Dunser, M. A Global Perspective on Acute Respiratory Distress Syndrome and the Truth about Hypoxia in Resource-limited Settings. Am. J. Respir. Crit. Care Med. 5–7 (2016).
6.	WHO. Global Pulse Oximetry Project. (2008).
7.	Funk, L., Weiser, T. & Berry, W. Global operating theatre distribution and pulse oximetry supply: an estimation from reported data. Lancet 1055–1061 (2010).
8.	Bagchi, S. Telemedicine in rural India. PLoS Med e82 (2006).
9.	Varma, N. Role of the surface electrocardiogram in developing countries. J. Electrocardiol. 612–614 (2010).
10. Silva, I. & Moody, G. An Open-source Toolbox for Analysing and Processing PhysioNet Databases in MATLAB and Octave. J. Open Res. Softw. 2, (2014).
11. Costa, M., Moody, G., Henry, I. & Goldberger, A. L. PhysioNet: an NIH research resource for complex signals. J. Electrocardiol. 36, 139–144 (2003).
Appendix A
Item	Dimensions	Cost (USD)
Stethoscope head	44.30mm x 62.45mm x 17.80mm	$0.206
Stethoscope Y piece	70.89mm x 29.94mm x 9.00mm	$0.09
Stethoscope ear tube	170.79mm x 83.62mm x 5.80mm	$0.07
Stethoscope ring	r = 21mm, h = 7mm	$0.012
Stethoscope spring	91.25mm x 111.62mm x 15.05mm	$0.032
Silicone tubing	40 cm – 12mm OD, 8mm ID / 2x 9 cm – 6mm OD, 4 mm ID	$1.93
Diaphragm	r = 20mm	$0.06
Ear plugs with mold	n/a	$0.43
Total		$2.83

Table 1. Bill of materials for the Glia model stethoscope (100% infill)

Calibration and comparison of 3D printed Glia model stethoscopes to the gold standard Littmann Cardiology III. Stethoscope output responses were measured using the equipment setup described in the methods. Each stethoscope model recorded input sound at multiple frequencies and the change in amplitude between input and recorded sound was documented for each stethoscope. Further comparison between the Glia Model 2 and the gold standard Littmann Cardiology III is shown in with the absolute Δmagnitude plotted above.
Appendix B
Item	Cost (CAD)
Infrared Emitter 940nm	0.38
640nm LED	0.77
25x Header and Wire House Connector	4.00
Light To Frequency & Light To Voltage Converter	2.06
FFC & FPC Connectors	1.44
ARM Micrrocontrollers	1.96
Fixed Inductors WE-TPC 2811	2.07
Speakers & Transducers	2.71
Voltage Regulators	0.90
2x Tactile Switches	0.67
3x Thick Film Resistor 1/20watt 1Mohms	1.13
Thick Film Resistor 1/10watt 590Kohms	0.15
Thick Film Resistor 1/10watt 2.0Kohms	0.15
LDO Voltage Regulator	0.58
4x Multilayer Ceramic Capacitors – MLCC SMD/SMT 0603 0.1uF 16V	0.58
2x Multilayer Ceramic Capacitors MLCC – SMD/SMT 6.3V 10uF	0.55
2x Multilayer Ceramic Capacitors MLCC – SMD/SMT 0603 4.7uF 6.3V	0.32
9x Multilayer Ceramic Capacitors MLCC – SMD/SMT 0603 1uF 10V	1.31
LCD w/Touch Panel	4.7
Total	26.43

Table 2. Bill of Materials for the Glia Pulse Oximeter