Tourniquet

Stethoscope

Section B – Project descriptions
242. What scientific or technological uncertainties did you attempt to overcome? (Maximum 350 words)

A tourniquet is considered to be an essential medical device in pre-hospital and emergency medical settings. A tourniquet is used to reduce or prevent blood loss from an extremity, and is typically used in patients who have suffered a penetrating trauma. Despite the ubiquity of tourniquets in the developed world, in the developing world tourniquets are often not widely available. Tragically, hemorrhagic deaths from limb exsanguination are common in conflict regions due to barriers access to this life-saving medical device. Tourniquets are scarce in underdeveloped regions such as the Gaza Strip. This scarcity can be attributed to both the high cost-per-unit of the tourniquet in addition to resource blockades which prevent the entry of medical supplies. For example, commercially available Combat Application Tourniquets (CAT) cost approximately USD$50 per piece. Moreover, the CAT is best suited for use in adults and do not work on a pediatric population, whereas the Gaza population is mostly children.

It was uncertain if an effective, low-cost, open-access tourniquet could be designed for mixed adult and pediatric use. No validated, open-access tourniquet was currently available at the time of project initiation.

244. What work did you perform in the tax year to overcome the scientific or technological uncertainties described in line 242?
(Summarize the systematic investigation or search) (Maximum 700 words)

The hypothesis of this project was that an effective, low-cost, open-access tourniquet could be designed, constructed and validated for mixed adult and pediatric use.

In this section, we will discuss the work that was done in this tax year on the design, construction and validation of a low-cost tourniquet.

Testing of varied tourniquet prototypes and materials testing was conducted to develop an effective tourniquet design. In laboratory testing of a preliminary prototype of the tourniquet on manikins, it was discovered that twisting the windlass too many times resulted in a catastrophic failure of the tourniquet. Consequently, corrections to the tourniquet design were implemented. The new model was then tested to failure and it was found to withstand approximately 5 times more force as the original design before non-catastrophic failure resulting in an unstable tourniquet.

 A prototype of Glia’s 3D printed tourniquet could be manufactured in Gaza at a cost of 25 NIS (USD$7). It was tested in the lab and in initial field trials with success.

This updated tourniquet design was produced and deployed in large quantity for field testing during violent Israeli suppression of the protests along the Gaza border.
Ten tourniquets were deployed on patients wounded with live ammunition meeting criteria for tourniquet use. Three tourniquets failed. Seven tourniquets operated as intended.

Field testing the tourniquet in large quantity revealed both the strengths and weaknesses of the device for use in trauma situations. Paramedics, first-responders and clinicians who used the device reported difficulties opening the packaging – identifying an area for improvement in future iterations of the device. Additionally, a buckle failed on a moderately critical patient while in the first phase of application while using the latest model of tourniquet which was an unexpected occurrence.

In conclusion, this project saw the first large deployment of open source, 3D-printed tourniquets in a battlefield setting. The tourniquets were well-received by patients, bystanders and first responders. Field testing large numbers of tourniquets on trauma patients revealed several moderate and critical errors in the device and identified areas for design improvement.

246 What scientific or technological advancements did you achieve or attempt to achieve as a result of the work described in line 244? (Maximum 350 words)

The technological advancement achieved by this project is confirmation that a novel, low-cost, 3D printed tourniquet could be designed, constructed and tested to demonstrate efficacy for use in trauma patients. Laboratory testing and field testing of various prototypical tourniquets allowed for the identification of design flaws in the tourniquet prototypes. This research project allowed for the development of an effective and operational ‘Glia’ model tourniquet.